Abbott i-STAT Alinity Point of Care Analyser

This Ward Handbook has been designed to ensure that the i-STAT Alinity System is being used within the governance guidelines of the LTHT Point of Care Policy.

Consumables

i-STAT Creatinine Cartridges REF 03P84-25

i-STAT Controls Level 1 REF 06F12-01

i-STAT Controls Level 3 REF 06F14-01

All consumables should be ordered directly from Abbott via your local supplies department, using the reference numbers provided above.

Please note: All required stock, including items needed by POCT for IQC verification purposes, are to be purchased by the Radiology department.

Contact Details

The Point of Care Testing (POCT) Team can be contacted between the hours of Mon-Sun 08:30-17:00 on:

Ext: 22338 (LGI) or 64791 (SJUH)

Mobile: 07775996028

Email: leedsth-tr.pointofcare@nhs.net

Training

All users must attend a training session every 2 years.

Training sessions can be arranged directly with the clinical area Key trainer who can be identified by contacting POCT or can be arranged directly with POCT. Key trainers must be trained by POCT.

For further information, see the i-STAT Alinity User Guide [POCT-SOP-15] and Operator Manual [POCT-MAN-15].

Documentation

Please complete the sheets within this Ward Handbook, compliance will be audited every 12 months.

Electronic Simulator Record - Performed every 24 hours and recorded

Quality Control Record - 2 levels performed weekly and recorded

Sample Record - Every sample or EQA result recorded

Reagent Register - Product information for new cartridge deliveries

Clew Updates - Every 6 months the iSTAT will require a CLEW software update. The analyser will give a warning TWO WEEKS before CLEW expiry. The POCT team will arrange for the CLEW update to be performed.

Electronic Simulator Record

The electronic simulator needs to be performed every 24 hours with the results recorded below. This should also be performed when a new lot of cartridges are received.

Analyser serial number __________________ *This field MUST be filled in*

| Date | Simulator number | Pass/Fail (If fail – action) | Operator |
| --- | --- | --- | --- |

Quality Control Record

Both level 1 and 3 liquid QC samples must be performed every week and when a new lot number of cartridges are received. All results must be recorded below.

** IMPORTANT- QC PREPARATION **

- QC ampoules must be left at room temperature for 30 minutes. - Shake the QC ampoule vigorously for 10 seconds before opening. - Snap the tip of the ampoule at the neck immediately before use. - QC testing must be performed within 10 minutes of opening. Failure to follow these preparation steps may lead to incorrect QC results. **QC results MUST be checked against the stated range in this document. Operator must then select on the meter if the QC has passed or failed** Please enter all IQC data into the SharePoint live document titled “Radiology Creatinine IQC”. Alternatively, QC results can be sent to the Point of Care team weekly. leedsth-tr.pointofcare@nhs.net

| QC Level | Lot No. | Exp | Min |  | Max | Units |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | 101194 | 28/02/27 | 363 | - | 394 | µmol/L |
| 3 | 121194 | 28/02/27 | 44 | - | 48 | µmol/L |

Analyser serial number __________________ *This field MUST be filled in*

| Date | Cartridge lot number | QC Lot Number | QC Level | Result (µmol/L) | Pass/Fail | Operator |
| --- | --- | --- | --- | --- | --- | --- |
| 02/07/20 | A20099 | 101123 | 1 | 400 | F (repeat) | Emma B |
| 02/07/20 | A20099 | 101123 | 1 | 390 | P | Emma B |

Sample Record

When using the meter to run a patient or EQA sample all the information must be recorded below.

| Date | Cartridge |  | Patient ID (Name and DOB plus NHS/PAS/CRIS) | Results |  | i-STAT Operator | Transcription checked by |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Lot NO. | Expiry Date |  | Creatinine µmol/L | eGFR |  |  |

Reagent Register

| Date of Receipt | Cartridge/QC | Lot Number | Expiry date | Temperature Correct? | Operator |
| --- | --- | --- | --- | --- | --- |

Please record all the details below on the receipt of any new cartridges or QC solutions.